## Sheet1
Part # | Vendor | Description | QTY | Unit Price (EUR) | Price (CHF) | Assembly | Subassembly | Comment
Camera
01-IRIS-15-PCIe-M-16-C | Photometrics | IRIS-15-PCIe sCMOS camera | 1 | 12350 | 12350 | Detection arm |  | FOV 25 mm diagonal
| Hamamatsu | Orca Lightning | 0 | 25000 | 0 |  |  | Alternative. FOV 29 mm diagonal
Excitation laser combiner options
Recommended laser combiner specs: 3-4 lasers: (405 optional), 488, 561, 638 nm, power 50-150 mW. Digital and analog power modulation. Coupling to SM fiber with FC/APC end connector.
| Oxxius | L4Cc 488-561-638 nm, 100 mW each, 2 fiber ports, switching module MDL-FSTM | 1 | 24000 | 24000 | Excitation |  | Recommended. Includes switching module between 2 laser ports, controlled by TTL signal.
1008062 | FISBA | FISBA READY Beam Bio 488-520-638nm | 0 | 7400 | 0 |  |  | Low-cost alternative, but no 561 nm line possible. Buggy software, often has troubles to start. Clunky wiring interface, need to make your own BNC breakouts.
| Cobolt | Cobolt Skyra 488-561-638, power 50-100 mW/ch | 0 | 20000 | 0 |  |  | Alternative. Moderate-cost 4-channel laser, with mature software and good support. Requires switching mechanism or laser duplication.
Sample & Microscope Stages
Dual-LS100-FTP | ASI | Two parallel LS-100 servo actuators | 1 | 5254 | 5254 | Sample motion |  | Sample Y stage
LS-1007-K | ASI | Pair of right angle brackets allow LS units to be mounted vertically 68mm x 48mm x 154mm | 1 | 520 | 520 | Sample motion
LS-50-AMCCH | ASI | LS-50 Linear Stage, with sub-micron accuracy, and 50 mm of travel | 1 | 1807 | 1807 | Sample motion |  | Sample X stage
LS-5012 | ASI | Plate for attaching flat surface of linear stage to breadboard. 3.9" x 5.9" | 1 | 132 | 132 | Detection arm | Detection path
mesoSPIM-Bridge | ASI | Bridge part | 1 | 250 | 250 | Sample motion |  | Bridge for XZR stage assembly
LS-50-AMCLH | ASI | LS-50 Linear Stage, with sub-micron accuracy, and 50 mm of travel | 1 | 1807 | 1807 | Detection arm | Detection path | Focus (F) stage, moves camera and detection lens assembly
DV-6010-C | ASI | Dovetail mount pair (VTS-2112 & CDZ-1010) 60 x 62 x 10 mm overall dimensions. | 1 | 200 | 200 | Sample motion |  | Cross-coupling X and Z stage adapter
LS-50S-AMECH | ASI | LS-50 Linear Stage, with sub-micron accuracy, 50 mm of travel, and stainless steel slide | 1 | 1939 | 1939 | Sample motion |  | Sample Z-stage
C60-3060-CMR-1011 | ASI | LS-50 slide to C-mount adapter | 1 | 110 | 110 | Sample motion |  | Rotary stage adapter
C60-3060-SRS | ASI | Rotary stage with collet for holding a sample. | 1 | 2450 | 2450 | Sample control |  | Sample Rotation (R) stage
TG8-BASIC | ASI | Controller: Includes TG8, TGCOM, 2x TGDCM2, SA-JOY+ZF Contols four DC servo motors, | 1 | 4800 | 4800 | Motion control |  | Motion controller with pre-installed cards
TGDCM2 | ASI | Dual Axis motion control card - one slot | 1 | 870 | 870 | Motion control |  | Motion controller card
Sample handling
KB25/M | Thorlabs | Complete 25 mm x 25 mm Kinematic Base, Top and Bottom Plates, M4 Counterbores | 2 | 70 | 140 | Sample handling |  | Magnetic bases to attach sample holders and immersion cuvette
KBB25/M | Thorlabs | Bottom Plate Only of the KB25/M Kinematic Base, M4 Counterbore | 6 | 36 | 216 | Sample handling |  | Sample holder side (nests, not balls) of the kinematic base
RLA150/M | Thorlabs | Dovetail Optical Rail, 150 mm, Metric | 1 | 43 | 43 | Sample handling |  | Immersion cuvette base
RC2/M | Thorlabs | Dovetail Rail Carrier, 50.8 mm x 25.4 mm, M6 Counterbore, M4 Taps | 1 | 29 | 29 | Sample handling |  | Rail carrier for immersion cuvette base
RD-PDT-S | Radiant Dyes | Prism Turn Table Small | 1 | 195 | 195 | Sample handling |  | Base for immersion cuvette
RD-PDT-S-to-RC2-height40mm.ipt | 3D printed | Adapter for mating Radyan Dyes prism turn table to Thorlabs rail carrier RC2 | 1 | 10 | 10 |  |  | See file RD-PDT-S-to-RC2-height40mm.ipt (.stl) for details.
Computer
Recommendations: at least 64 GB RAM, 1 TB SSD drive (for image processing)
|  | Imaging Workstation with Screen | 1 | 8000 | 8000
Excitation arms
33-494 | Optotune | Optotune EL-16-40-TC-VIS-5D-1-C lenses | 2 | 780 | 1560 | Excitation | ETL Unit
88-940 | Optotune | Optotune Lens Driver EL-E-4i | 2 | 280 | 560 | Excitation | ETL Unit | Have to be modified: BNC cables need to be soldered
88-941 | Optotune | Optotune 6-way Hirose Male/Female Cable, 1m Length | 2 | 85 | 170 | Excitation | ETL Unit
SM1L10 | Thorlabs | SM1 Lens Tube, 1.00" Thread Depth, One Retaining Ring Included | 2 | 13 | 26 | Excitation | ETL Unit
SM1M20 | Thorlabs | SM1 Lens Tube Without External Threads, 2" Long, Two Retaining Rings Included | 2 | 17 | 34 | Excitation | ETL Unit
SM1A10 | Thorlabs | Adapter with External C-Mount Threads and Internal SM1 Threads, 4.4 mm Spacer | 2 | 18 | 36 | Excitation | ETL Unit
varies | Digitec or elsewhere | Powered USB-Hub | 1 | 30 | 30 | Excitation | ETL Unit | Simple USB-Hub for powering the ETL-Drivers
91863 | Nikon | Nikon AF-S 50 mm f/1.4 G Objective | 2 | 400 | 800 | Excitation | Excitation obj | Have to be modified
SM1FCA | Thorlabs | FC/APC Fiber Adapter Plate with External SM1 (1.035"-40) Threads, Wide Key (2.2 mm) | 2 | 31 | 62 |  | Laser input | One per arm
CP33/M | Thorlabs | SM1-Threaded 30 mm Cage Plate, 0.35" Thick, 2 Retaining Rings, M4 Tap | 10 | 16 | 160 |  | Cage assembly | 5 per arm
ER8-P4 | Thorlabs | Cage Assembly Rod, 8" Long, Ø6 mm, 4 Pack | 2 | 41 | 82 |  | Cage assembly | 8 rods per arm
AC254-060-A-ML | Thorlabs | f=60 mm, Ø1" Achromatic Doublet, SM1-Threaded Mount, ARC: 400-700 nm | 2 | 99 | 198 |  | Laser collimator | 1 per arm, if laser combiner has 2 fiber outputs (Oxxius)
KCB1E/M | Thorlabs | Right-Angle Kinematic Elliptical Mirror Mount with Tapped Cage Rod Holes, 30 mm Cage System and SM1 Compatible, M4 and M6 Mounting Holes | 4 | 189 | 756 |  | Laser folding | M1, M2 fold mirrors, in each arm
BBE1-E02 | Thorlabs | 1" Broadband Dielectric Elliptical Mirror, 400 - 750 nm | 4 | 110 | 440 |  | Laser folding | M1, M2 fold mirrors, in each arm
PH150/M | Thorlabs | Ø12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=150 mm | 4 | 13 | 52 |  | Mechanical support | Optical cage support, 2 per arm
PH75/M | Thorlabs | Ø12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=75 mm | 2 | 8 | 16 |  | Excitation obj mount
PH40/M | Thorlabs | Ø12.7 mm Post Holder, Spring-Loaded Hex-Locking Thumbscrew, L=40 mm | 4 | 7 | 28 |  | Mechanical support | Optical cage support, 2 per arm
BA1S/M | Thorlabs | Mounting Base, 25 mm x 58 mm x 10 mm | 6 | 5 | 30 |  | Mechanical support | Clamping post holders to optical table
TR100/M | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 100 mm | 4 | 6 | 24 |  | Mechanical support | Optical cage support, 2 per arm
TR40/M | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 40 mm | 4 | 5 | 20 |  | Mechanical support | Optical cage support, 2 per arm
TR75/M | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 75 mm | 2 | 5 | 10 |  | Excitation obj mount
XRN25P | Thorlabs | Compact 25 mm Travel Linear Translation Stage, Side Micrometer, 1/4"-20 Taps | 2 | 398 | 796 |  | M3 mirror assembly | Linear stage for M3 mirror adjustment
mirror-bracket.stl | 3D printed | Braket for M3 kinematic mirror mount | 2 | 1 | 2 |  | M3 mirror assembly | Bracket for M3 mirror mounting on linear stage. Folder /custom-parts/excitation-mirror-bracket/
XRN-B1/M | Thorlabs | Base Plate for Stages with 2" Dovetails, Metric Slot Spacing | 2 | 30 | 60 |  | M3 mirror assembly | Adapter plate for linear stage
LCP03/M | Thorlabs | 60 mm Blank Cage Plate, M4 Tap | 2 | 36 | 72 |  | Galvo assembly | Needs modification, see LCP03M_modified.ipt
Deprecated 10mm Galvos
GVS211/M | Thorlabs | 1D Large Beam (10 mm) Diameter Galvo System, Broadband Mirror (-E02), Metric, Power Supply Not Included | 0 | 1757 | 0 |  | Galvo assembly | Galvo cables must be ordered long (80 cm), and controller boards must be tuned accordingly, ask Thorlabs. Short (default) cables enforce placing the controller boards inside the housing, and they generate a lot of heat.
Galvo bracket must be modified (using milling machine), see Galvo-bracket-modified.ipt
GCE001 | Thorlabs | Galvo Driver Card Cover | 0 | 59 | 0 |  | Galvo assembly
GPS011-EC | Thorlabs | 1D or 2D Galvo System Linear Power Supply, 230 VAC | 0 | 495 | 0 |  | Galvo assembly | Power supply for 2 galvos. Voltage and plugs are country specific!
To be tested:
QS15X-AG | Thorlabs | Ø15 mm Beam Galvo System, X-Axis Protected Silver Mirror | 2 | 2104 | 4208 |  | Galvo assembly | Standard cables are 60 cm long. Sufficient?
CBLS3F | Thorlabs | Command and Power Cables for QS15/20/30/45, SS, & SP Series Galvo Scanners | 2 | 82 | 164 |  | Galvo assembly
GPWR15 | Thorlabs | ±15 V Power Supply for BLINK & QS, SS, SP, & XG Series Galvo Scanners | 1 | 390 | 390 |  | Galvo assembly | Power supply for 2 galvos. Voltage and plugs are country specific!
ER4-P4 | Thorlabs | Cage Assembly Rod, 4" Long, Ø6 mm, 4 Pack | 2 | 25 | 50 |  | Galvo assembly
LCP06/M | Thorlabs | 60 mm Cage Plate with Ø2" Double-Bore Optic Mount, M4 Tap | 2 | 35 | 70 |  | Excitation obj mount | Needs modification (using lathe or mill), see LCP06M-Nikon-modified.ipt | Todo: design 3D-printable alternative
SM2NFM | Thorlabs | Adapter with External SM2 Threads and Nikon Female F-Mount Ring | 2 | 92 | 184 |  | Excitation obj mount | The F-mount steel ring must be taken out and mounted into the modified LCP06M plate using screws.
Cage60mm-2sides-InnerD75mm-L75mm.stl | 3D printed | Laser safety snap-on covers | 2 | 1 | 2 |  | Excitation obj mount | See files files in folder /custom-parts/laser-safety/
Detection: Camera and lens mount
LS-50-AMCLH | ASI | LS-50 Linear Stage, with sub-micron accuracy, and 50 mm of travel | 1 |  |  |  |  | see Sample and Microscopy Stages
XT66SD-250 | Thorlabs | 6 mm Single Dovetail Rail with Mounting Counterbores, L = 250 mm | 1 | 63 | 63
XT66P2/M | Thorlabs | Rail Carriage for 66 mm Rails with M4 & M6 Taps | 2 | 69 | 138
Lensagon1.2x-XT66P2M.ipt | Custom (3D print) | Mounting clamp (for Lensagon 1.2x, T25M-12-155I) | 1 | 10 | 10 |  |  | 3D printed (2 parts) and bolted together with M6. Modify the bore diameter in IPT file to fit other lenses.
LS-5012 | ASI | Plate for attaching flat surface of linear stage to breadboard. 3.9" x 5.9" | 1 |  |  |  |  | see Sample and Microscopy Stages
Detection: Filter wheels (OPTIONAL)
Option 1: Servo-driven filter wheel. ONLY for telecentric lenses or long-WD objectives (> 75 mm)
5 positions available. Motorized filterwheel is useful when multiple bandpass filters are used during an acquisition. This helps separate emission spectra of the labels.
184-902-0064-000 | Mouser | Robotis Dynamixel MX-28R | 1 | 250 | 250 | Detection path | Filter wheel servo
184-903-0163-100 | Mouser | Robotis FR07-S101K Set | 1 | 13 | 13 | Detection path | Filter wheel servo
895-USBRS485WE5000BT | Mouser | USB to RS485 Serial Converter Cable | 1 | 43 | 43 | Detection path | Filter wheel servo
FilterWheel-2inch-5pos.stl | 3D printed | 5-Position filter wheel for 2-inch filters | 1 |  |  |  |  | See folder /custom-parts/filter-wheel/
Option 2: Motorized flip mount ONLY for telecentric lenses or long-WD objectives (> 75 mm) |  | Todo: Adapt the code and test this option.
Two positions: (1) quadband emission filter (2) no filter, alignment mode. Fits most of the use cases.
MFF102/M | Thorlabs | Motorized Filter Flip Mount with Ø2" Optic Holder, M4 Tap | 0 | 656 | 0
TR20/M | Thorlabs | Ø12.7 mm Optical Post, SS, M4 Setscrew, M6 Tap, L = 20 mm | 0 | 4.65 | 0
BA1S/M | Thorlabs | Mounting Base, 25 mm x 58 mm x 10 mm | 0 | 4.87 | 0
Filters
Filters have to be selected according to lasers & needs (flurescent markers and dyes used)
F57-405SG | AHF | QuadLine Rejectionband ZET405/488/561/640 in 50 mm diameter | 1 | 1075 | 1075 | Detection path |  | A frequently used quad-band filter, allows all fluorescence through
F37-520SG | AHF | 520/35 BrightLine HC in 50 mm diameter | 0 | 1350 | 0 | Detection path |  | Optional
F37-542SG | AHF | 542/27 BrightLine HC in 50 mm diameter | 0 | 1350 | 0 | Detection path |  | Optional
F37-515SG | AHF | 515/LP BrightLine HC Longpass-Filter in 50 mm diameter | 0 | 1445 | 0 | Detection path |  | Optional
F76-561SG | AHF | 561 LP Edge Basic Longpass-Filter in 50 mm diameter | 0 | 1130 | 0 | Detection path |  | Optional
F76-594SG | AHF | 594 LP Edge Basic Longpass-Filter in 50 mm diameter | 0 | 1130 | 0 | Detection path |  | Optional
Detection: Lenses
11041 | Lensation | 1.2x Lensagon T25M-12-155I | 1 | 6734 | 6734 | Detection path | Camera assembly | 1.2x telecentric lens, F-mount, suitable for sensors up to 25 megapixel and with an image format of 32mm (diagonal).
IN PROGRESS
Detection: Other components
Electronics / Signal generation
778512-01 | National Instruments | NI PXI-6733 with 8 16-Bit Waveform Analog Outputs | 1 | 3582 | 3582 | Signal generation |  | The signal generation card, 8 analog and 8 digital channels.
781163-01 | National Instruments | NI PXIe-1073 Integrated MXIe, 5 Periph Slots | 1 | 1630 | 1630 | Signal generation |  | Chassis to mount the signal generation card
779504-01 | National Instruments | NI PCIe-8361, 1 Port MXI-Express Interface | 1 | 886 | 886 | Signal generation |  | Connects the Chassis to the PC
779500-03 | National Instruments | MXI-Express Cable, Gen 1 x1, Copper, 3m | 1 | 163 | 163 | Signal generation |  | Cable between the Chassis and the PC
763065-01 | National Instruments | Power Cord, 220V, 10A (plug country-specific) | 1 | 11 | 11 | Signal generation
777643-01 | National Instruments | BNC-2110 Noise Rejecting, Shielded BNC Connector Block | 1 | 520 | 520 | Signal generation
199006-01 | National Instruments | SH68-68-EPM Shielded Cable, 68 D-Type to 68 D-Type, 1M | 1 | 193 | 193 | Signal generation
779632-01 | National Instruments | PXI-6259 (32 Analog Inputs, 48 Digital I/O, 4 Analog Outputs) | 0 | 2800 | 0 | Signal generation |  | OPTIONAL. Only if you need more than 4 laser lines.
PAA236R | Thorlabs | SMB Coaxial Cable, 90° SMB Female to BNC Male, 36" (914 mm) | 10 | 12 | 120 | Signal generation |  | Depending on the laser control interface (Oxxius: SMB).
CA2806 | Thorlabs | SMA Coaxial Cable, SMA Male to BNC Male, 6" (152 mm) | 1 | 14 | 14 | Signal generation |  | Depends on laser interface. Oxxius: IO connector must be modified, pin24 (MDL-FSTM) and pin16 (GND Digital), to accept arm switching signal.
2249-C-120 | Thorlabs | 0.3 m BNC Cable | 4 | 15.25 | 61 | Signal generation
2249-C-24 | Thorlabs | 0.6 m BNC Cable | 10 | 16.5 | 165 | Signal generation
2249-C-60 | Thorlabs | 1.5 m BNC cable | 10 | 24 | 240 | Signal generation
2249-C-120 | Thorlabs | 3m BNC cable | 6 | 24 | 144 | Signal generation
Alignment Tools
SI254 | Thorlabs | Shearing Interferometer with a 10-25.4 mm Beam Diameter Shear Plate | 1 | 488 | 488
SI100P | Thorlabs | Shear Plate, 5-10 mm Beam Diameter | 1 | 184 | 184
SICP | Thorlabs | 30 mm Cage System Adapter for Shearing Interferometers with Diffuser Viewing Screen | 1 | 76 | 76
CCM1-G01/M | Thorlabs | 30 mm Cage Cube-Mounted Protected Aluminum Turning Mirror, M4 Tap | 1 | 184 | 184
AC254-100-A-ML | Thorlabs | f=100 mm Ø1" Achromatic Doublet SM1-Threaded Mount ARC: 400-700 nm | 1 | 94.1802 | 94.1802
SM1V15 | Thorlabs | Ø1" Adjustable Lens Tube, 1.31" Travel Range | 1 | 36.8 | 36.8
SM1T2 | Thorlabs | SM1 (1.035"-40) Coupler, External Threads, 0.5" Long | 2 | 21.21 | 42.42
DGM05 | Thorlabs | Digital Indicator | 1 | 304.75 | 304.75 |  |  | Very useful
DIGC6 | Thorlabs | Digital Calipers | 1 | 148.21 | 148.21 |  |  | Very useful
R2L2S3P2 | Thorlabs | Grid Distortion Target, 1.5" x 1.5", 250 µm Grid Spacing | 1 | 348.45 | 348.45 |  |  | Very useful
FFM1 | Thorlabs | 30-mm-Cage-Compatible Rectangular Filter Mount | 1 | 60.95 | 60.95
DG10-1500-H1-MD | Thorlabs | Ø1" SM1-Mounted Frosted Glass Alignment Disk w/Ø1 mm Hole | 1 | 34.5 | 34.5 |  |  | Nice to have
DG20-1500-H2-MD | Thorlabs | Ø2" SM2-Mounted Frosted Glass Alignment Disk w/Ø2 mm Hole | 1 | 51.45 | 51.45 |  |  | Nice to have
CPA1 | Thorlabs | 30 mm Cage System Alignment Plate with Ø1 mm Hole | 2 | 13.8 | 27.6
LCPA1 | Thorlabs | 60 mm Cage Alignment Plate | 1 | 19.09 | 19.09
LCP01/M | Thorlabs | 60 mm Cage Plate, SM2 Threads, 0.5" Thick, M4 Tap (Two SM2RR Retaining Rings Included) | 2 | 40.51 | 81.02 |  |  | Nice to have as spares
CP20S | Thorlabs | 30 mm Cage System Iris Ø20.0 mm Maximum Aperture | 2 | 83.5981 | 167.1962 |  |  | Nice to have
SM1D12D | Thorlabs | SM1 Ring-Actuated Iris Diaphragm (Ø0.8 - Ø12 mm) | 4 | 70.104 | 280.416 |  |  | Nice to have
G061008000 | Qioptiq | Mounting Plate Without Hole | 2 | 28.75 | 57.5 |  |  | Have to be modified with SM1 threads
G061372000 | Qioptiq | Carrier FLR 40-20 | 2 | 38.8728 | 77.7456
| Fisso | Strato M-28 movable arm | 1 | 205 | 205 |  |  | Very useful
| Fisso | Thread adapter female M8/ male M6 thread reduction for Fisso | 2 | 12 | 24 |  |  | Very useful
General lab supplies
LAB-16 | Newport | Ultra Jet Canned Air Duster, 10 oz. (300ml) | 4 | 16.5 | 66 |  |  | Or other canned duster
19569/16OZ | Cargille | Immersion oil from cargille for Clarity: Code 50350 nD=1.4587 fused silica matching oil | 2 | 140 | 280 |  |  | A decent supply is recommend (though it is not consumed much)
MC-50E | Thorlabs | Lens Tissues, 25 Sheets per Booklet, 50 Booklets in a Closeable Box | 1 | 85.1 | 85.1
BKF12 | Thorlabs | Black Aluminium Foil | 2 | 32 | 64 |  |  | Sometimes useful
T743-2.0 | Thorlabs | High-Performance Black Masking Tape, 2" x 180' (50 mm x 55 m) Roll | 2 | 51.58 | 103.16
PSS7 | Thorlabs | 7-Piece Precision Screwdriver Set | 1 | 60 | 60 |  |  | Or other, cheaper sets
HW-KIT6/M | Thorlabs | Imperial Screw Thread Adapter Kit | 1 | 136.48 | 136.48
HW-KIT6/M | Thorlabs | Metric Screw Thread Adapter Kit | 1 | 136.48 | 136.48
HW-KIT1/M | Thorlabs | M4 Cap Screw and Hardware Kit | 1 | 60.95 | 60.95
HW-KIT3/M | Thorlabs | M4 Setscrew and Hardware Kit | 1 | 119.416 | 119.416
HW-KIT4/M | Thorlabs | M6 Setscrew and Hardware Kit | 1 | 105.8 | 105.8
HW-KIT2/M | Thorlabs | M6 Cap Screw and Hardware Kit | 2 | 119.6 | 239.2
HW-KIT5 | Thorlabs | 4-40 Cap Screw and Hardware Kit | 1 | 100 | 100
TC3/M | Thorlabs | 15-Piece Balldriver and Hex Key Kit with Stand, Metric | 2 | 92 | 184
TC2 | Thorlabs | 20-Piece Balldriver and Hex Key Kit with Stand, Imperial | 1 | 96.23 | 96.23
SPW801 | Thorlabs | Adjustable Spanner Wrench | 1 | 102.44 | 102.44
SPW602 | Thorlabs | SM1 Spanner Wrench, Graduated, Length = 3.88" | 2 | 28 | 56
|  | Multiport-power outlets | 4
|  | Cable wraps | 1 |  |  |  |  | Lots of them
|  | Isopropanol p.a. for lens cleaning | 1
Custom components
| Mechanical Workshop | Galvomount-22mm-diameter-Thorlabs-Left.ipt | 1 |  |  | Excitation Path |  | Only if Thorlabs Scanners are used | Aluminium
| Mechanical Workshop | Galvomount-22mm-diameter-Thorlabs-Right.ipt | 1 |  |  | Excitation Path |  | Only if Thorlabs Scanners are used | Aluminium
| Mechanical Workshop | 60mm-blank-cageplate-LCP03M-Nikon-MOD.ipt | 2 |  |  | Excitation Path |  | Modified LCP03M | Aluminium
| Mechanical Workshop | MVX-10-Bracket-V5.ipt | 1 |  |  | Detection Path |  |  | Aluminium
| Mechanical Workshop | Large-cuvette-mount-30mm-V1.ipt | 0 |  |  | Sample Mounts |  | if 30 mm cuvettes are desired | POM (Polyoxymethylen)
| Mechanical Workshop | Large-cuvette-mount-40mm-V3.ipt | 2 |  |  | Sample Mounts |  | Order one for each planned immersion cuvette mount / imaging medium | POM (Polyoxymethylen)
| Mechanical Workshop | Large-cuvette-mount-50mm-V3.ipt | 2 |  |  | Sample Mounts |  | Order one for each planned immersion cuvette mount / imaging medium | POM (Polyoxymethylen)
| Mechanical Workshop | M-060-Adapter.ipt | 1 |  |  | Sample Mounts |  |  | Aluminium
Will be redesigned | Mechanical Workshop | Sample-cuvette-clamp-20mm-V3.ipt | 10 |  |  | Sample Mounts |  | Redesign is being worked on | POM (Polyoxymethylen)
Sample cuvettes | Material: glass, RI=1.52. For quartz (e.g. CLARITY imaging, RI=1.47), replace UG qwith UQ in the part name.
| Portmann Instruments | UG-205-H80 10x20x80 | 0 | 518.5 | 0 |  |  | Whole-Mouse CNS Cuvettes
| Portmann Instruments | UG-753-H100 40x40x100 | 0 | 764 | 0
| Portmann Instruments | UG-205-H120 10x20x120 | 0 | 816 | 0 |  |  | Whole-Mouse CNS Cuvettes
| Portmann Instruments | UG-753-H120 40x40x120 | 0 | 952 | 0
| Portmann Instruments | UG-751 20x20 mm | 0 | 99.5 | 0
| Portmann Instruments | UG-752 30x30 | 0 | 128.5 | 0
| Portmann Instruments | UG-753 40x40 | 0 | 156 | 0
| Portmann Instruments | UG-754 50x50 | 0 | 184.5 | 0
| Portmann Instruments | UG-203 5x10x45 | 0 | 106.5 | 0
| Portmann Instruments | UG-204 10x10x45 | 6 | 106.5 | 639
| Portmann Instruments | UG-205 10x20x45 | 10 | 121.5 | 1215
|  |  |  | SUM | 99111.534